Gyrocomp Gyratory Compactor - Lockout technical specifications

The gyratory compactor shall comply with relevant ASTM, AASHTO, SuperPave and EN standards: EN 12697-10, EN 12697-31, ASTM D6925, AASHTO T312 (TP4), SHRP M002. It shall be capable of producing samples from laboratory prepared or plant mix hot mix asphalt, warm-mix asphalt, foamed mix and cold mixes, including mixes containing emulsified binder. It shall be constructed with a highly rigid steel frame to ensure excellent angle control and accuracy. The essential components of the gyratory compactor shall include the following specification requirements:

  - Approved as Superpave Gyratory Compactor in United States

  - Load applied by pneumatic piston, adjustable from 80 to 800 kPa on 150 mm dia.  specimen and 160 to 1400 kPa on 100 mm dia. specimen.

  - Internal angle of gyration adjustable from 0.7 to 1.4°, standard factory set: 0.82° (EN12697-31) or 1.16° (ASTM D6925, AASHTO T312, SHRP M-002)

  - Number of cycles adjustable from 1 to 999.

  - Gyration rate shall be adjustable from 5 to 60 cycles/min (30 cycles/min required by standard)

  - Height measurement with 0.1 mm accuracy

  - Gyratory compactor shall be capable of compacting the specimen with three possible termination modes: Number of gyrations, Maximum Height and Density.

  - Capable of compacting cold mixes utilising bituminous emulsions;

  - Capable of fabricating specimens of compacted size: 100 mm & 150 mm dia. and height up to 200 mm for both dia.

  - Safety switch for protection

  - Capable of operating at 230 V, 50 Hz or 110 V 60 Hz

  - User interface both with touch screen panel display and remote control from PC software

  - Load measurement by pressure transducer

  - The compaction in a new mould and the specimen extrusion from the previous mould can proceed simultaneously for productivity

  - Sliding transparent door with safety interlock

Optional:

  - Electromechanical operated specimen extruder capable of extruding both 150 mm and 100 mm diameter and 200 high samples with integrated bench.

  - Availability of calibration kit for

    - Height: It should be supplied with standard height block,

    - Pressure: Vertical force testing device complete with official calibration certificate for testing and verifying load.

    - Gyration rate

  - Compaction moulds, bottom plates and piston face:

    - Moulds with 150 mm and 100 mm internal diameter and height of 250 mm

    - Moulds shall be smooth of 1.60 μm in accordance with ASME B46.1

    - Moulds shall have steel walls that are at least 7.5 mm thick and hardened to Rockwell C48 or better.

    - Mould plates and ram heads are to be fabricated from steel with a minimum Rockwell hardness of C48.

    - Piston Face for 100 diameter mould and 150 mm diameter mould

    - Mould weight for easy lifting less or equal to:

      - 10 Kg for 150 mm diameter mould

      - 7 Kg for 100 mm diameter mould

  - Perforated moulds of 100 & 150 mm dia. to run tests on cold emulsified asphalt mix

  - Calibration kit for Angle of gyration: Gyratory internal angle measurer should also be quoted for calibration of angle as per EN, ASTM and AASHTO standards:

  - The angle calibrator should perform top and bottom angle measurement as specified by standard, the average should be considered as internal angle of machine.

  - Data measured should be downloadable via RS232 cable at the end of measurement.

  - Angle of Gyration accuracy: 0.01°

  - Calibration techniques / procedure shall fulfil ASTM, EN and AASHTO requirements

  - Software:

- User-friendly Windows based PC software for remote test control and test data acquisition and display of results in real time including:

      - Specimen height

      - Number of gyrations

      - Density

      - Air voids

Inputs:

      - Weight

      - Maximum theoretical density

- Data elaboration of single and multiple test with facility to perform Superpave Mix Design procedure